INTERNATIONAL ELECTROTECHNICAL COMMISSION SYSTEM FOR CERTIFICATION TO STANDARDS RELATING TO EQUIPMENT FOR USE IN EXPLOSIVE ATMOSPHERES (IECEx SYSTEM)

TITLE: Re-assessment Report for the continued acceptance of Physikalisch-Technische Bundesanstalt (PTB) as an Accepted Ex Test Laboratory (ExTL), within the IECEx System, Equipment Scheme 02.

Circulation to: Members of the IECEx Management Committee, ExMC

INTRODUCTION

In accordance with the 5 year re-assessment plan for the surveillance and monitoring of bodies within the IECEx System, the following document contains the IECEx Re-assessment Report for the continued acceptance of Physikalisch-Technische Bundesanstalt (PTB) as an Accepted Ex Test Laboratory (ExTL).

This report is hereby submitted for endorsement during the ExMC 2015 Meeting being held in Christchurch, New Zealand.

Chris Agius

IECEx Secretariat

| IECEx Secretariat Australia Square Level 33, 264 George Street Sydney NSW 2000 Australia | Tel: +61 2 4628 4690 Fax: +61 2 4625 3480 Email: chris.agius@iecex.com |
| --- | --- |

IECEx ASSESSMENT REPORT FOR

Physikalisch-Technische Bundesanstalt (PTB)

IECEx TEST LABORATORY, ExTL

Type of Assessment: (please mark)

Initial assessment for Candidate ExTL

Re-Assessment of ExTL				X

Scope Extension of ExTL

Surveillance Audit

- OBJECT AND FIELD OF APPLICATION

  - Country:

Germany

  - Name of Candidate TL

Physikalisch-Technische Bundesanstalt (PTB)

  - Members of the Assessment Team

Heinz Berger – IECEx Officer - IECEx Lead Assessor

Ajay Maira - IECEx Expert Assessor

  - Place and Date of Assessment

Bundesallee 100

38116 Braunschweig, Germany

September 3rd to 5th, 2014

  - Assessment References

- IECEx 02 Equipment Scheme Rules (current version)

- IECEx OD/003 Assessment Procedures (current version)

- IECEx OD 005-1, and -2; Manufacturer Assessment (current version)

- IECEx OD/009 Equipment Scheme Procedures(current version)

- IECEx OD/18  Checklist 17025 (current version)

- IECEx OD/24 Witness Testing/manufacturer and users Facility

- ISO/IEC 17025:2005

- IECEx Technical Capability Document (TCD)

- ExTAG decision sheets (DSs)

- OD's related to technical issues

  - Scope of Application

| Number | Title | Acceptance |
| --- | --- | --- |
| 60079-0 Edition 3.1 Edition 4 Edition 5 Edition 6 | Explosive atmospheres - Part 0: Equipment - General requirements | YES |
| 60079-1 Edition 3.2 Edition 4 Edition 5 Edition 6 Edition 7 | Explosive atmospheres - Part 1: Equipment protection by flameproof enclosures 'd' | YES |
| 60079-2 Edition 4 Edition 5 | Explosive atmospheres - Part 2: Equipment protection by pressurized enclosures 'p' | YES |
| 60079-5 Edition 2 Edition 3 | Explosive atmospheres - Part 5: Equipment protection by powder filling 'q' | YES |
| 60079-6 Edition 2 Edition 3 | Explosive atmospheres - Part 6: Equipment protection by oil immersion 'o' | YES |
| 60079-7 Edition 3 Edition 4 | Explosive atmospheres - Part 7: Equipment protection by increased safety 'e' | YES |
| 60079-11 Edition 4 Edition 5 Edition 6 | Explosive atmospheres - Part 11: Equipment protection by intrinsic safety 'i' | YES |
| 60079-15 Edition 1 Edition 2 Edition 3 Edition 4 | Explosive atmospheres - Part 15: Equipment protection by type of protection 'n' | YES |
| 60079-18 Edition 1 Edition 2 Edition 3 | Explosive atmospheres - Part 18: Equipment protection by encapsulation "m" | YES |
| 60079-25 Edition 1 Edition 2 | Explosive atmospheres - Part 25: Intrinsically safe systems | YES |
| 60079-26 Edition 1 Edition 2 | Explosive atmospheres - Part 26: Equipment with equipment protection level (EPL) Ga | YES |
| 60079-27 Edition 1 Edition 2 | Explosive atmospheres - Part 27: Fieldbus intrinsically safe concept (FISCO) | YES |
| 60079-28 Edition 1 | Explosive atmospheres - Part 28: Protection of equipment and transmission systems using optical radiation | YES |
| 60079-30-1 Edition 1 | Explosive atmospheres - Part 30-1: Electrical resistance trace heating - General and testing requirements | YES |
| 60079-31 Edition 1 Edition 2 | Explosive atmospheres - Part 31: Equipment dust ignition protection by enclosure "t" | YES |
| 61241-0 Edition 1 | Electrical apparatus for use in the presence of combustible dust - Part 0: General requirements | YES |
| 61241-1 Edition 1 | Electrical apparatus for use in the presence of combustible dust - Part 1: Protection by enclosures 'tD' | YES |
| 61241-1-1 Edition 2 | Electrical apparatus for use in the presence of combustible dust - Part 1: Electrical apparatus protected by enclosures and surface temperature limitation - Specification for apparatus | YES |
| 61241-4 Edition 1 | Electrical apparatus for use in the presence of combustible dust - Part 4: Type of protection 'pD' | YES |
| 61241-11 Edition 1 | Electrical apparatus for use in the presence of combustible dust - Part 11: Protection by intrinsic safety 'iD' | YES |
| 61241-18 Edition 1 | Electrical apparatus for use in the presence of combustible dust - Part 18: Protection by encapsulation 'mD' | YES |

  - Candidate TL Persons Interviewed

| Name | Position |
| --- | --- |
| Dr. Ulrich Johannsmeyer | Head of Sector Explosion Protection of the PTB Conformity Assessment Body, Head of Department 3.6 "Explosion Protection in Sensor Technology and Instrumentation" |
| Dipl.-Ing. Reinhard Wilkens | Quality assurance representative Explosion Protection |
| Dr. Uwe Klausmeyer | Head of ExTL and of Department 3.5 "Explosion Protection in Energy Technology” |
| Thomas Horn | Head of Working Group 3.61 “Intrinsic Safety” |
| Mario Graube | Head of Working Group 3.64 “Production Control in Explosion Protection” |
| Uwe Völkel | Project Engineer |
| Axel Linne | Project Engineer |
| Thomas Uehlken | Scientific Engineer |
| Rüdiger Strutz | Technician |
| Dr. Christian Lehrmann | Project Engineer |

  - Legal Entity of the TL

PTB is a Governmental Body of the Federal Republic of Germany belonging to the Ministry of Economics and Energy, based on the law about metrology (EinhZeitG, BGBI. I S. 408, Article 4, clause 68 of August 7, 2013).

  - Associated ExCB

The ExTL is integral with the ExCB. The ExTL is incorporated in PTB Division 3 with the following departments involved in IECEx activities: 3.7 (mechanical and electrostatics), 3.5, 3.6, (both electrical).

  - Financial Support

PTB has the financial stability and resources required for the operation of the IECEx CB through the governmental system according to the financial regulation of the German law, # 0913. PTB has in income from testing and certification activities in the IECEx System. PTB is providing organization, technical and material support for the Proficiency Testing Program.

  - History

PTB is the highest technical authority under the auspices of the Federal Ministry of Economics and Energy (BMWI) with the tasks: fundamentals of metrology, metrology for economy, metrology for society, international affairs. As national metrology institute, it is providing scientific and technical services and stands for progress and reliability in metrology for the benefit of society, trade and industry, and science.

PTB was founded in 1887 as the Physikalisch-Technische Reichsanstalt on the basis of the ideas of Werner von Siemens and Hermann von Helmholtz and on their joint initiative. PTB has grown to about 1900 staff members. The traditional site of the PTB is in Berlin-Charlottenburg (since 1887). Since 1947, the PTB has been residing in Braunschweig. About 1300 employees in eight divisions work on an area of 1 km2; the Presidential Board of the PTB has its seat in Braunschweig.

- ORGANISATION

  - Names, Titles and Experience of the Senior Executives

| Name | Title | Experience |
| --- | --- | --- |
| Uwe Klausmeyer | Head of ExTL and of Department. 3.5 "Explosion Protection in Energy Technology” | 30 years in Ex |
| Michael Beyer | Head of Department “Basics of Explosion Protection” | 25 years in Ex |
| Ulrich Johannsmeyer | Head of Sector Explosion Protection of the PTB Conformity Assessment Body, Head of Department 3.6 "Explosion Protection in Sensor Technology and Instrumentation" | 39 years in Ex |
| Frank Lienesch | Head of Working Group 3.52 "Explosion-protected Electrical Drive Systems" | 20 years in Ex |
| Thomas Horn | Head of Working Group 3.61 “Intrinsic Safety” | 15 years in Ex |
| Detlef Markus | Head of Working Group 3.51 “Flame Transmission Processes” | 17 years in Ex |
| Jörg Flötotto | Head of Working Group 3.53 “Explosion-protected Electrical Power Equipment” | 15 years in Ex |
| Mario Graube | Head of Working Group 3.64 “Production Control in Explosion Protection” | 23 years in Ex |
| Martin Thedens | Head of Working Group 3.73 “Physical Ignition Processes” | 19 years in Ex |
| Reinhard Wilkens | Quality assurance representative | 30 years in Ex |

  - Name, Title and Experience of the Quality Management Representative

| Name | Title | Experience |
| --- | --- | --- |
| Dr. Jens Könemann | Quality responsible person of Division 3 |  |
| Reinhard Wilkens | Quality assurance representative | 30 years in Ex |
| Dr. Monika Schumann | Assistant Quality assurance representative | 17 years in Ex |

  - Name and Title of Nominated Principal Contact

| Name | Title | Comments |
| --- | --- | --- |
| Uwe Klausmeyer | Head of ExTL and of Department. 3.5 "Explosion Protection in Energy Technology | uwe.klausmeyer@ptb.de |

  - Employees

| Name | Title | Comments/Exp. In Ex |
| --- | --- | --- |
| Gerlinde Haas | Head of administrative office | 25 years in Ex |
| Tanja Eckert | Member of administrative office | 11 years in Ex |
| Uwe Dreger | Project engineer of ExTL | Department 3.5; 23 years |
| Wolfgang Ellwardt | Project engineer of ExTL | Department 3.6; 29 years |
| Thomas Horn | Project engineer of ExTL | Department 3.6; 15 years |
| Jörg Flötotto | Project engineer of ExTL | Department 3.5; 15 years |
| Manfred Glander | Project engineer of ExTL | Department 3.5; 15 years |
| Mario Graube | Project engineer of ExTL | Department 3.6; 23 years |
| Kurt Hassel | Project engineer of ExTL | Department 3.6; 25 years |
| Martin Krämer | Project engineer of ExTL | Department 3.6; 27 years |
| Rainer Kulessa | Project engineer of ExTL | Department 3.6; 23 years |
| Christian Lehrmann | Project engineer of ExTL | Department 3.5; 12 years |
| Frank Lienesch | Project engineer of ExTL | Department 3.5; 20 years |
| Axel Linne | Project engineer of ExTL | Department 3.6; 19 years |
| Frank Neugebauer | Project engineer of ExTL | Department 3.5; 20 years |
| Andreas Piotrowski | Project engineer of ExTL | Department 3.5; 20 years |
| Jürgen Rüberg | Project engineer of ExTL | Department 3.5; 34 years |
| Monika Schumann | Project engineer of ExTL | Department 3.5; 17 years |
| Martin Thedens | Project engineer of ExTL | Department 3.7; 19 years |
| Uwe Völkel | Project engineer of ExTL | Department 3.5; 23 years |
| Reinhard Wilkens | Project engineer of ExTL | Department 3.6; 30 years |
| Ernst Petereit | Project engineer of ExTL | Department 3.5; 20 years |
| Marc Shields | Technician | Department 3.5; 13 years |
| Rüdiger Strutz | Technician | Department 3.5; 34 years |
| Lutz Schubert | Technician | Department 3.6; 25 years |
| Adam Kübler | Technician | Department 3.6; 30 years |
| Kathrin Hermann | Technician | Department 3.6; 15 years |

  - Organizational Structure

See attached organisation chart in ANNEX 1a and ANNEX 1b

or http://www.ptb.de/cms/en/fachabteilungen/abt3.html

- RESOURCES

There are 39 employees involved in the testing activity. The test facilities cover the requirements of the IECEx Scheme.

- DOCUMENTATION

  - Quality Manual

The KBS (Conformity Assessment Body) of PTB is maintaining a comprehensive quality manual.

The quality documentation system of the certification activities is integrated in the overall quality system of PTB. It consists of two main levels where for each level a quality manual is available with the corresponding procedures. The conformity assessment body belongs to level 2. IECEx is described in level 2. Specific procedures, e.g. work instructions, can be found in the documentation of sector 1 of the KBS (explosion protection).

The whole quality system of PTB with all levels and sections was presented during the assessment by use of the Intranet. Specific documents concerning the IECEx System were reviewed in detail and found to meet the requirements of IECEx.

  - Procedures

In level 1 general procedures are described. Additionally special procedures of the ExTL are part of the documentation of the quality management system of division 3 in the intranet. The documentation meets the requirements of the IECEx System.

  - Work Instructions

In level 2 the specific procedures and requirements for the ExTLs (explosion protection) are described. The work instructions of the ExTLs are part of the intranet quality documentation of division 3.

  - Records

Records are stored according to PTB Procedure number QM-VA-09 and Quality Manual of the KBS, section 2.5.

ExTR extracts, QAR extracts and CoCs are stored electronically on the IECEx server. Other records such as complete ExTRs and QARs as well as drawings are retained hard copy. The storage time is 30 years.

  - Document Change Control

According to PTB QM procedure number VA-08.

  - Test Records

Records are handled/stored according to PTB Procedure number QM-VA-09 and Quality Manual of the KBS, section 2.5.

ExTR extracts, QAR extracts and CoCs are stored electronically on the IECEx server. Other records such as complete ExTRs and QARs as well as drawings are retained complete in hard copy. The storage time is 30 years.

- TEST REPORTS

  - Test Reports Written and Reviewed

Number of test reports issued under national schemes and IECEx in the preceding four years for each type of protection:

| Standards | Title | Number of test reports |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | 2011 | 2012 | 2013 | 2014 | Total |
| 60079-0 and 61241-0 | Explosive atmospheres - Part 0: Equipment - General requirements | --- | --- | --- | --- | Parts 0 included in numbers below |
| 60079-1 | Explosive atmospheres - Part 1: Equipment protection by flameproof enclosures 'd' | 57 | 94 | 49 | 30 | 230 |
| 60079-2 | Explosive atmospheres - Part 2: Equipment protection by pressurized enclosures 'p' | 5 | --- | --- | 2 | 7 |
| 60079-5 | Explosive atmospheres - Part 5: Equipment protection by powder filling 'q' | 2 | 4 | --- | 2 | 8 |
| 60079-6 | Explosive atmospheres - Part 6: Equipment protection by oil immersion 'o' | --- | --- | --- | --- | --- |
| 60079-7 | Explosive atmospheres - Part 7: Equipment protection by increased safety 'e' | 109 | 161 | 62 | 34 | 366 |
| 60079-11 | Explosive atmospheres - Part 11: Equipment protection by intrinsic safety 'i' | 58 | 100 | 57 | 32 | 247 |
| 60079-15 | Explosive atmospheres - Part 15: Equipment protection by type of protection 'n' | 13 | 12 | 9 | 7 | 41 |
| 60079-18 | Explosive atmospheres - Part 18: Equipment protection by encapsulation "m" | 11 | 38 | 10 | 7 | 66 |
| 60079-25 | Explosive atmospheres - Part 25: Intrinsically safe systems | 1 | 1 | 1 | --- | 3 |
| 60079-26 | Explosive atmospheres - Part 26: Equipment with equipment protection level (EPL) Ga | 18 | 15 | 12 | 4 | 49 |
| 60079-27 | Explosive atmospheres - Part 27: Fieldbus intrinsically safe concept (FISCO) | 2 | 1 | --- | --- | 3 |
| 60079-28 | Explosive atmospheres - Part 28: Protection of equipment and transmission systems using optical radiation | 1 | 5 | 1 | 2 | 9 |
| 60079-30-1 | Explosive atmospheres - Part 30-1: Electrical resistance trace heating - General and testing requirements | --- | 4 | --- | --- | 4 |
| 60079-31 | Explosive atmospheres - Part 31: Equipment dust ignition protection by enclosure "t" | 54 | 81 | 48 | 27 | 210 |
| 61241-0 | Electrical apparatus for use in the presence of combustible dust - Part 0: General requirements | --- | --- | --- | --- | See comment above |
| 61241-1 | Electrical apparatus for use in the presence of combustible dust - Part 1: Protection by enclosures 'tD' | 24 | 41 | --- | 1 | 66 |
| 61241-1-1 | Electrical apparatus for use in the presence of combustible dust - Part 1: Electrical apparatus protected by enclosures and surface temperature limitation - Specification for apparatus | --- | 1 | --- | --- | 1 |
| 61241-4 | Electrical apparatus for use in the presence of combustible dust - Part 4: Type of protection 'pD' | --- | --- | --- | --- | --- |
| 61241-11 | Electrical apparatus for use in the presence of combustible dust - Part 11: Protection by intrinsic safety 'iD' | 15 | 20 | --- | 11 | 46 |
| 61241-18 | Electrical apparatus for use in the presence of combustible dust - Part 18: Protection by encapsulation 'mD' | --- | --- | --- | --- | --- |

- CALIBRATION

PTB is the National Metrology Institute of Germany and hence there are significant facilities available.  However, there is a lot of calibration done within the explosion protection departments. Hence, equipment for use in the Ex area is calibrated internally and calibration certificates are available. External calibration is used as well. A number of certificates were reviewed and found to meet the requirements of the IECEx. PTB uses a database to control calibration information.

- CONFIDENTIALITY

According to quality manual of the KBS, section 4.2 and PTB procedure number QM-VA-05. Several CVs were reviewed concerning the confidentiality declaration. All CVs checked contained a signed declaration.

In order to follow the new requirements concerning impartiality as described in ISO/IEC 17065, an updated confidentiality declaration form will be introduced and signed by the employees.

- NATIONAL ACCREDITATION

The PTB Ex-laboratories were accredited by the German Accreditation Service ZLS according to ISO/IEC 17025 under the registration number ZLS-P-698/06 until July 2011. Since then, ZLS was no longer able to issue accreditation certificates. Hence, PTB was not in the position to present a valid accreditation certificate. However, PTB is in the process to make the decision concerning a national accreditation (DAkkS). Therefore, the one year rule applies to perform a surveillance assessment in case PTB is not in the position to present an accreditation document for the ExTL.

- RECOGNITION AND AGREEMENTS

According to Quality Manual of the KBS, section 4.7, based on Mutual Recognition Agreements, (MRA), Memorandums of Understanding (MOU) and Letters of Intent, PTB has agreements with organizations in the following countries:

USA, Canada, Japan, China, Ukraine, Brazil, Australia, Korea, Germany and Switzerland.

PTB received a letter of acceptance from the commander of the United States Coast Guard for the IEC standards 60079-0, 1, 2, 5, 6, 7, 11, 15 und 18.

- INTERNAL AUDIT AND PERIODIC REVIEW

According to PTB Procedures number QM-VA-02 and QM-VA-01.

The audit program for 2014 was presented.  The audit was performed on 14.4.2014. The requirements of the IECEx are fulfilled.

The last management review was performed on 29.7.2014 and the report was available during the assessment.

The overall management review chaired by the president of PTB is scheduled for the 21.10.2014.

- COMPLAINTS AND APPEALS (Including appeals to IECEx)

Complaints and appeals are handled according to Quality Manual of the KBS, section 4.4 and PTB Procedure number QM-VA-14.

Complaints and appeals are subject to discussion during the yearly management review and are reported regularly to the quality manager of PTB.

No complaints and appeals were received in the past 5 years concerning the IECEx Scheme.

- SPECIAL FACTS TO BE NOTED

  - Supporting Documentation

Copies of additional supporting information for this assessment have been provided to the audited organization and the IECEx Secretariat. These include:

- Details of issues raised and how these have been resolved

- Checklist for ISO/IEC 17025

- Completed Technical Capability Document (TCD)

- Photos of the facilities

  - Tests Witnessed or Competence checked

The Following tests were successfully witnessed by the assessment team:

| Standard: IEC 60079-0 Ed 6 General Requirements |  |  |
| --- | --- | --- |
| Clause | Requirement – Test |  |
| 26.4.2 | Resistance to impact |  |
| 26.4.3 | Drop test |  |
| 26.4.5 | Degree of protection (IP) by enclosures – dust test |  |
| 26.4.5 | Degree of protection (IP) by enclosures – water test |  |
| 26.5.1 | Thermal tests |  |
| 26.5.2 | Thermal shock test |  |
| 26.5.3 | Small component ignition test (Group I and Group II) |  |
| 26.6 | Torque test for bushings |  |
| 26.8 | Thermal endurance to heat |  |
| 26.9 | Thermal endurance to cold |  |
| 26.12 | Earth continuity |  |
| 26.13 | Surface resistance test of parts of parts of enclosures of non-metallic materials |  |
| 26.14 | Measurement of capacitance |  |
| 26.15 | Verification of ratings of ventilating fans |  |
| 26.16 | Alternative qualification of elastomeric sealing O-rings |  |
| A.3.1 | Tests of clamping of non-armoured and braided cables |  |
| A.3.2 | Tests of clamping of armoured cables |  |

| Standard: IEC 60079-1 Ed 6 Flameproof enclosure ‘d’ |  |  |
| --- | --- | --- |
| Clause | Requirement – Test |  |
| 15.1.2 | Determination of Reference Pressure |  |
| 15.1.3 | Overpressure test |  |
| 15.2 | Test for non-transmission of an internal ignition |  |
| 15.5 | Tests of flameproof enclosures with breathing and draining devices |  |
| B.1.2 | Sintered metal elements - bubble test pore size |  |
| B.1.3 | Sintered metal elements - Density |  |
| B.1.4 | Sintered metal elements - Open porosity and/or fluid permeability |  |
| C.3.1 | Cable glands - Sealing test |  |
| C.3.3.1 C.3.4.1 | Type tests for Ex blanking elements - Torque test |  |
|  | Availability and adequacy of equipment |  |

| Standard: IEC 60079-2 Ed 5 Pressurized enclosure ‘p’ |  |  |
| --- | --- | --- |
| Clause | Requirement – Test |  |
| 16.1 | Maximum overpressure test |  |
| 16.2 16.2.1 16.2.2 | Leakage test |  |
| 16.3 16.3.1 | Purging test, protective gas is air |  |
|  | Availability and adequacy of equipment |  |
| 16.3 16.3.2 | Purging test, protective gas is inert |  |
| 16.4 16.4.2.1/2 16.4.3.1/2 16.4.4.1/2 | Purging test and dilution test, internal release |  |
| 16.5 | Minimum overpressure |  |
| 16.6 16.6.1 | infallible containment system, Overpressure test |  |
| 16.6 16.6.2 | Infallible containment system, Infallibility test |  |
| 16.7 | Overpressure test for a containment system with a limited release |  |
| 16.8 | Verifying ability of the pressurized enclosure to limit internal pressure |  |

| Standard: IEC 60079-5: 2007 Ed 3 Powder filling ‘q’ |  |  |
| --- | --- | --- |
| Clause | Requirement – Test |  |
| 5.1.1 | Pressure type test of enclosure |  |
| 5.1.2 | Verification of the degree of protection of the enclosure |  |
| 5.1.3 | Flammability of materials |  |
| 5.1.4 | Dielectric strength of the filling material |  |
| 5.1.5 | Maximum temperatures |  |

| Standard: IEC 60079-6: 2007 Ed 3 Oil immersion "o" |  |  |
| --- | --- | --- |
| Clause | Requirement – Test |  |
| 5.1.1 | Overpressure test on sealed enclosures |  |
| 5.1.2 | Reduced pressure test on sealed enclosures |  |
| 5.1.3 | Overpressure test on unsealed enclosures |  |

| Standard: IEC 60079-7 Ed 4 Increased safety |  |  |  |
| --- | --- | --- | --- |
| Clause |  | Requirement – Test |  |
| 6.1 |  | Dielectric Tests |  |
| 6.2.1 |  | tE of cage rotor machines |  |
| 6.2.3 |  | Additional tests for machines liable to ‘sparking’ |  |
| 6.3.1 |  | Mechanical tests for screw lamp holders other than E10 |  |
| 6.3.2 |  | Abnormal operation of luminaries with tubular fluorescent lamps |  |
| 6.3.3 |  | Sulphur dioxide test for the connection of bi-pin lamp caps to lamp holders |  |
| 6.3.4 |  | Vibration test for luminaries with bi-pin lamps |  |
| 6.4 |  | Measuring instruments and instrument transformers – thermal and inter-turn test |  |
| 6.5 |  | Transformers other than instrument transformers - thermal tests |  |
| 6.6.2 |  | Batteries – insulation test |  |
| 6.6.3 |  | Batteries – mechanical shock test |  |
| 6.6.4 |  | Ventilation of battery container |  |
| 6.7 |  | General purpose junction boxes – heating test |  |
| 6.8 |  | Resistance heating devices (not trace heating) |  |
| 6.9 |  | Terminal insulating material tests – thermal conditioning then pull test |  |
| 6.2.1 | tE of cage rotor machines |  |  |

| Standard: IEC 60079-11 Ed 5 Intrinsic Safety |  |  |
| --- | --- | --- |
| Clause | Requirement – Test |  |
| 10.1 | Spark Testing |  |
| 10.2 | Temperature Tests |  |
| 10.3 | Dielectric Strength Tests |  |
| 10.5 | Tests for Cells and Batteries |  |
| 10.6 | Mechanical Tests |  |
| 10.7 | Tests for apparatus containing piezoelectric devices |  |
| 10.8 | Type tests for diode safety barriers and safety shunts |  |
| 10.9 | Cable Pull tests |  |
| 10.10 | Transformer tests |  |
| 10.11 | Optical isolators tests |  |

| Standard: IEC 60079-15 Ed 4 Type of protection ‘n’ |  |  |
| --- | --- | --- |
| Clause | Requirement – Test |  |
| 22.3.1 | Thermal endurance |  |
| 22.4 | Enclose break tests |  |
| 22.5 | Sealed devices test |  |
| 22.6 | Restricted breathing test |  |
| 22.7 | Screw lamp holder tests |  |
| 22.12 | Insulation resistance test for batteries |  |
| 22.13 | Ignition tests for large or high-voltage machines |  |

| Standard: IEC 60079-18 Ed 3 Type of protection ‘m’ |  |  |
| --- | --- | --- |
| Clause | Requirement – Test |  |
| 8.1.1 | Water absorption test |  |
| 8.1.2 | Dielectric strength tests on compound |  |
| 8.2.2 | Maximum temperature |  |
| 8.2.3.1 | Thermal endurance to heat |  |
| 8.2.3.2 | Thermal endurance to cold |  |
| 8.2.4 | Dielectric strength tests - circuits |  |
| 8.2.5 | Cable pull tests |  |
| 8.2.6 | Pressure tests |  |
| 8.2.7 | Endurance test for resettable devices with contacts |  |
| 8.2.7 | Endurance test for resettable devices without contacts |  |
| 8.2.8 | Sealing tests |  |

| Standard: IEC 60079-26 Ed 2 Type of protection ‘Ga’ |  |  |
| --- | --- | --- |
| Clause | Requirement – Test |  |
| General |  |  |

| Standard: IEC 60079-28 Ed 1 Part 28: Protection of equipment and transmission systems using optical radiation |  |  |
| --- | --- | --- |
| Clause | Requirement – Test |  |
| 1.1.1 | Optic power cell |  |

| Standard: IEC 60079-30-1 Ed 1 Part 30.1: Electrical resistance trace heating—General and testing requirements |  |  |
| --- | --- | --- |
| Clause | Requirement – Test |  |
| General |  |  |

| Standard: IEC 60079-31 Equipment dust ignition protection by enclosure "t" |  |  |
| --- | --- | --- |
| Clause | Requirement – Test |  |
| 6.1.1 | Dust exclusion by enclosures (IP 5X or 6X) with 2kPa or 4kPa depression |  |
| 6.1.2 | Thermal tests (including dust layer for ta) |  |
| 6.1.3 | Pressure test (4kPa or 2kPa) |  |

- COMMENTS (Including issues found during assessment)

The Vice-President of PTB, Dr. Roman Schwartz and the CEO of the PTB Conformity Assessment Organization, Dr. Harry Stolz, were present during the Re-Assessment.

Issues were found concerning the accreditation, description in an Ex “d” COCs, calibration, test report and test circuits. All open issues are resolved to the satisfaction of the assessment team.

- RECOMMENDATION

Based on the Re-assessment performed from September 3rd to 5th, 2014, the ExTL of PTB is recommended for continued acceptance in the IECEx Scheme as an IECEx Testing Laboratory (ExTL) according to the scope of the standards listed in this document.

| Lead Assessor Heinz Berger | Expert Assessor Ajay Maira |  |
| --- | --- | --- |

Date:  September 5th, 2014

List of Annexes:

ANNEX 1a:  Overall Organization Chart

ANNEX 1b:  Organization Chart of ExCB and ExTL

ANNEX 1a:

ANNEX 1b: