- ExTAG DS 2002/001A

2007 09 25

RECORD OF IECEx / ExTAG DECISIONS

| Standard: | Clause: 5.4, 7.1, 7.4 | Sheet No. 2002/001A |
| --- | --- | --- |
| IEC 60079-11:1999-02 |  |  |
| Subject: | Key words: | Decision of meeting |
| Electrochemical sensors | - Gas detectors | Date: 2002-10-21(original) Revised: 2007-09-25 to take into account editorial changes |
| Status of document: adopted | - Electrochemical sensors - Voltage generation | Originator of proposal: AU/BKI/DMT TC/SC involved: SC31G Revised by: Mr W Lawrence (FM) |
| Question: Can electrochemical sensors be considered as simple apparatus? Answer: Simple apparatus are: “Devices, in which, according to the manufacturer's specifications, none of the values 1.5 V, 0. 1 A, or 25 mW is exceeded." Electrochemical sensors generate electrical energy depending on the concentration of certain gases in air. The electrical output characteristics depend on concentration, temperature, humidity etc. Typical values under worst-case-condition have been reported*) to be: Gas open circuit load current Short circuit max. power Oxygen (0-100%) 0.9 V 10 mA 500 mA 112.5 mW (peaks; capacitive) Toxic gases 1.25 V 50 mA 1 A 300 mW (peaks; capacitive) Sometimes the maximum output values for sensors for toxic gases can only be reached with lethal concentrations. It is recommended to assess the characteristics only in the range of warning levels . In some cases it may be possible that the characteristics do not override the above mentioned limited. If the sensors are stand-alone-equipment they could be assessed to be simple apparatus But if they are integrated into an electronic circuitry with other voltage sources, the below mentioned rule of sheet No. 2002/005 shall be applied: When equipment contains more than one voltage source (for example, the main battery and a memory retention cell), then a check should be made that there is an infallible connection between one of the terminals of each source to form a common mode. If this is not the case then the possible addition of the voltage sources due to failure of tracks and components shall be considered. *) Paper presented by J. Hanko 1 BKI / HU |  |  |