- INTERNATIONAL ELECTROTECHNICAL COMMISSION (IEC) SYSTEM FOR CERTIFICATION TO STANDARDS RELATING TO EQUIPMENT FOR USE IN EXPLOSIVE ATMOSPHERES (IECEx SYSTEM)

TITLE: Draft Revision of DS 2015/016 Testing of thermocatalytic sensors with Ex ia IIC marking

- Circulated to: ExTAG – IECEx Testing and Assessment Group

INTRODUCTION

This document, is a proposed Revised version of ExTAG DS 2015/016 as prepared by Paul Kelly and Alexander Zalogin and has been prepared for consideration during the ExTAG Washington, USA, Meeting

Changes to the original edition are shown using the tracking tool.

Christine Kane

Julien Gauthier

Julien Gauthier

ExTAG Secretary

| Address: IECEx Secretariat Level 33 Australia Square 264 George Street Sydney NSW 2000 Australia Web: www.iecex.com | ExTAG Secretary Mr Julien Gauthier LCIE S.A. 33 Avenue du General Leclerc 92260 Fontenay-aux-Roses FRANCE Tel: +33 1 40 95 55 26 Fax: +33 1 40 95 89 37 Email : julien.gauthier@fr.bureauveritas.com |
| --- | --- |

COLLECTION OF IECEx / ExTAG DECISIONS

| Standard: IEC 60079-11:2011 Ed.6.0 | Clause: 5.6.2 | Date: (revised date) Originator of revised proposal: NANIO CCVE, UL LLC |
| --- | --- | --- |
| Subject: Testing of thermocatalytic sensors with Ex ia IIC marking Status of document: Approved | Key words: - thermocatalytic sensor - intrinsically safe circuit | TC/SC involved: IEC/TC 31 MT 60079-11, MT 60079-1 Draft Decision Sheet: ExTAG/401A/CD |
| Background: In clause 5.6.2 Temperature for small components for Group I and Group II of IEC 60079-11:2011 it is stated that: Requirements for temperatures of small components used in Group I or Group II equipment are provided in the small component temperature for Group I or Group II electrical equipment requirements of IEC 60079- 0 and the test requirements are provided in the small component ignition test of IEC 60079- 0. NOTE Where a catalytic or other chemical reaction can result specialist advice should be sought. According to the experience testing of thermocatalytic sensors with Ex ia IIC marking with a platinum sensing element without an enclosure in the homogeneous explosive testing mixture of between 22,5% and 23,5 % v/v diethyl ether and air shall not be considered as sufficient to confirm explosion safety, as such tests do not take into account possible thermocatalytic reactions. It’s shown by the experiments that under specific conditions heating of the sensing element in hydrogen-air mixture can cause an explosion even with switched off power. The specific experimental conditions that resulted in ignition involved: - Removal of the protective grid around the sensing element, - Flowing a mixture of 21,0% v/v hydrogen and air across the sensing element Ignition resulted under the above conditions both when the sensing element was under power, and when no power was applied. Note: No ignition resulted when flowing a mixture of 35,0% v/v hydrogen and air across the element. Also, no ignition resulted when the protective grid was not removed, even when a mixture of 21,0% v/v hydrogen and air was involved. Question: Is it sufficient to test thermocatalytic sensor (with a platinum sensing element intended to be used in gas analyzers, including those with flow boosters) with Ex ia IIC marking according to IEC 60079-0 as a small component in mixture of between 22,5% and 23,5 % v/v diethyl ether and air? Answer: First, it needs to be understood that there are no known reports of explosions in the field due to thermocatalytic reactions from gas analyzer sensing elements. Further, the issue of thermocatalytic reactions is only mentioned in IEC 60079-11 as an Informative Note. However, potential hazards associated with thermocatalytic reactions is a concern that should be considered. Thermocatalytic sensor with a platinum sensing element under high concentrations of explosive gas (including hydrogen/air) can ignite the explosive mixture of (21 ±2)% v/v hydrogen and air due to the development of autocatalysis reaction. Due to this concern, and due to the availability of type of protection “da” per IEC 60079-1:2014, use of type of protection “ia” should be avoided for IIC applications of such gas analyzer sensing elements. The sensors acting on the base of a thermocatalytic reaction should not be marked Ex ia IIC T4 Ga, but instead should be marked Ex da IIC based on a complaint design in accordance with IEC 60079-1:2014. It is recommended that both MT 60079-11 and MT 60079-1 address the issue of thermocatalytic reactions in their respective next edition efforts. |  |  |