- INTERNATIONAL ELECTROTECHNICAL COMMISSION (IEC) SYSTEM FOR CERTIFICATION TO STANDARDS RELATING TO EQUIPMENT FOR USE IN EXPLOSIVE ATMOSPHERES (IECEx SYSTEM)

- Title:  ExTAG/588/CD - Draft ExTAG Decision Sheet – Calibration and Operation of Climatic Chambers

- Circulated to: ExTAG – IECEx Testing and Assessment Group

INTRODUCTION

This document, ExTAG/588/CD Draft ExTAG Decision Sheet - Calibration and Operation of Climatic Chambers has been prepared by SGS BASEEFA, GB, as a result of discussion held during the ExTAG 2018 Cannes Meeting.

This document is issued for consideration during the EXTAG 2019 Dubai Meeting.

to

Christine Kane

Julien Gauthier

ExTAG Secretary

| Address: IECEx Secretariat Level 33 Australia Square 264 George Street Sydney NSW 2000 Australia Web: www.iecex.com | ExTAG Secretary Mr Julien Gauthier LCIE S.A. 33 Avenue du General Leclerc 92260 Fontenay-aux-Roses FRANCE Tel: +33 1 40 95 55 26 Fax: +33 1 40 95 89 37 Email : julien.gauthier@fr.bureauveritas.com |
| --- | --- |

COLLECTION OF IECEx / ExTAG DECISION

| Standard: IEC 60079-0 (Ed 6) IEC 60079-0 (Ed 7) | Clause: 26.8 26.8 | Draft Decision Sheet: ExTAG/588/CD |
| --- | --- | --- |
| Subject: Calibration and Operation of Climatic Chambers Status of document: DRAFT | Key words: - Temperature - Humidity | # Date: 2019/08/27 Originator of proposal: SGS Baseefa TC/SC involved: IEC/TC31 WG 22 |
| Background A questionnaire (ExTAG/510/Q) was circulated in July 2018 to all ExTLs requesting information on how they calibrated and used their Climatic (Humidity) Chambers for the tests of IEC 60079-0 clause 26.8. Just less than half the ExTLs responded and the responses indicated a wide variance in the methods used to calibrate the chambers. The results of the questionnaire formed the basis for a presentation and discussion at the ExTAG Training session held on 2018-09-18 at Cannes. During the presentation and discussion, it was particularly noted that at a nominal test condition of 95C and 90%RH, the permitted tolerance of ±2K on temperature was too wide, if the permitted tolerance of ±5%RH was not to be exceeded, since a change of 1K in the temperature, could result in a change of at least 4% in the Relative Humidity. At the formal ExTAG session fallowing the training session, the meeting accepted the need to issue a Decision Sheet and accepted in principle the content of this draft. Question Noting that TC31 WG22 thought it important to provide for testing with a climatic chamber setting of 90C 90%RH as an alternative to the previously specified 95C 90%RH, albeit with an increase in the exposure time by 50%, what is required to calibrate and use Climatic Chambers at the extreme conditions envisaged in IEC 60079-0 Clause 26.8? Answer In order to ensure comparability of results between testing laboratories, the following points should be taken into account: 1. Nine Point (or 15 point for over 2000 litre capacity) temperature calibration is to be conducted initially and at least every 3 years. (See IEC 60068-3-5.) 1.1 For 95C testing, the gradient should not exceed 2K from highest to lowest of the nine (15) points, otherwise the ±5% RH requirement throughout the chamber cannot be achieved. 1.2 For 90C testing (and lower) the gradient should not exceed 4K from highest to lowest of the nine (15) points. 2. The Climatic Chamber cannot be used for testing at a condition higher than the maximum temperature at which it is calibrated. 2.1 A chamber calibrated at both 80C and 90C can be used at any intermediate temperature. 3. TC31 WG22 introduced the alternative 90C test for 3 weeks to replace the 95C test for 2 weeks, to enable laboratories to use chambers that could not be calibrated at 95C. 3.1 It is preferable to use the 90C condition, the 95C condition having been retained in the standard primarily to cover pre-existing equipment. 4. Calibration of the chamber is considered to be an additional activity beyond the calibration of the chamber sensors connected to the chamber controller. A calibrated temperature and humidity transfer probe is to be sited at the centre of the chamber to get the comparison between the actual chamber conditions and the temperature and humidity indicated by the chamber controller. 4.1 As conditions within the chamber will be affected by the size and positioning of a test sample, the size of the sample must occupy less than 40% of the volume of the chamber, with the sample placed as close as possible to the centre of the chamber. 4.2 If it is necessary to test a sample greater than 40% of the volume of the chamber, additional calibration shall be undertaken with the actual sample or a dummy sample in position, particularly on both the upstream and downstream airflow sides of the sample. 4.3 Samples should never be located outside the designated working area of the chamber, i.e. the boundary of the eight corner mounted temperature sensors for the nine point method 4.4 The airflow from the fan(s) through the chamber must be constant through both calibration and use. Notes: 1. This Decision Sheet is subject to revision once current work on the same subject is concluded in IECEE. 2. The relevant standards are: IEC 60068-3-5 Confirmation of the performance of temperature chambers IEC 60068-3-6 Confirmation of the performance of temperature/humidity chambers IEC 60068-3-11 Calculation of uncertainty of conditions in climatic chambers (Part 3-11 is particularly useful in considering how to calibrate for, and test with, a chamber that is more than minimally loaded.) |  |  |