INTERNATIONAL ELECTROTECHNICAL COMMISSION SYSTEM FOR CERTIFICATION TO STANDARDS RELATING TO EQUIPMENT FOR USE IN EXPLOSIVE ATMOSPHERES (IECEx SYSTEM)

TITLE: Re-assessment and Scope Extension Report for the continued acceptance of DEKRA Testing and Certification GmbH, (DEKRA EXAM) an Accepted Certification Body (ExCB) in the IECEx Equipment Scheme 02, an Accepted Ex Testing Laboratory (ExTL) within the Equipment Scheme 02 as well as an Accepted ExCB for IECEx Certified Service Facilities Scheme, 03.

Circulation to: Members of the IECEx Management Committee, ExMC

INTRODUCTION

In accordance with the 5 year re-assessment plan for the surveillance and monitoring of bodies within the IECEx System, the following document contains the IECEx Re-assessment and Scope Extension Report for the continued acceptance of DEKRA Testing and Certification GmbH( DEKRA EXAM) an Accepted Certification Body (ExCB), in the IECEx Equipment Scheme 02, an Accepted Ex Testing Laboratory (ExTL) within the Equipment Scheme 02 as well as an Accepted ExCB for IECEx Certified Service Facilities Scheme, 03

During the re-assessment the IECEx Assessment Team took the opportunity to also assess, DEKRA Testing and Certification GmbH facilities, equipment and competence to undertake testing and certification to the Standards –

| IEC TS 60079-42 Edition 1.0 | Explosive atmospheres- Electrical safety devices for the control of potential ignition sources for Ex-Equipment | Scope extension |
| --- | --- | --- |

The report is hereby submitted for voting by the ExMC.

This document is hereby submitted for ExMC approval via correspondence using the IECEx on-line voting system.  ExMC Members are requested to submit their vote via the IECEx On-line Ballot System  by the closing date 2020 12 24

Please refer to OD 050 for guidance on the “IECEx On-line voting system.”

Chris Agius

IECEx Secretariat

| IECEx Secretariat Australia Square Level 33, 264 George Street Sydney NSW 2000 Australia | Tel: +61 2 4628 4690 Fax: +61 2 4625 3480 Email: info@iecex.com |
| --- | --- |

IEC System for certification to standards relating to equipment for use in Explosive Atmospheres (IECEx System)

IECEx Assessment Report Form

IECEx Assessment Report Form for use by IECEx Assessment Teams to report Assessments conducted according to the IECEx Assessment Procedures of

- Operational Document IECEx OD003-2 for the Certified Equipment Scheme

- Operational Document IECEx OD316-5 for the Certified Service Facility Scheme

- Operational Document IECEx OD422 for the IECEx Conformity Mark Licensing System

IECEx ExCB/ExTL assessment report for

DEKRA Testing and Certification GmbH

INTERNATIONAL
ELECTROTECHNICAL
COMMISSION

CONTENTS

- Assessment information	6

1.1	Type of Body covered by this assessment:	6

1.2	Type of assessment:	6

1.3	Details of body	6

1.3.1	Country	6

1.3.2	Name of body	6

1.3.3	Name and title of nominated principal contact	6

1.4	Assessment information	6

1.4.1	Members of the assessment team	6

1.4.2	Place(s) of assessment	6

1.4.3	Assessment date(s)	6

1.5	Application information and background information on the assessment	7

1.6	Scopes	7

1.6.1	ExCB scope for equipment certification scheme	7

1.6.2	ExTL scope	9

1.6.3	ExCB scope for Service Facilities Scheme	9

Common information	10

2.1	Legal entity of body	10

2.2	Financial support	10

2.3	History	10

2.4	Documentation	10

2.4.1	Quality manual	10

2.4.2	Procedures	10

2.4.3	Work instructions	11

2.4.4	Records (including test records where relevant)	11

2.4.5	Document change control	11

2.5	Confidentiality	11

2.6	Communication with public and customers (Hard copy and Electronic)	11

2.7	Recognitions and agreements	11

2.8	Internal audit	11

2.9	Management review	12

2.10	Contracting, subcontracting and witness testing	12

2.10.1	Contracting	12

2.10.2	Subcontracting	12

2.10.3	Witness testing	12

2.11	Training and competence	12

2.12	Complaints and appeals (including appeals to IECEx)	13

2.13	Impartiality	13

2.14	Commenting on ExTAG Documents	13

2.15	Special facts to be noted	13

2.16	Supporting documentation	13

2.17	Recommendations	13

ExCB for IECEx Certified Equipment Scheme	15

3.1	Assessment references	15

3.1.1	General references	15

3.1.2	Additional references applied for this assessment	15

3.2	Candidate ExCB persons interviewed	15

3.3	Associated ExTL(s)	15

3.4	Associated certification functions	15

3.5	National marks and certificates	15

3.6	Standards accepted	16

3.7	National differences to IEC standards	16

3.8	Organisation	16

3.8.1	Names, titles and experience of the senior executives	16

3.8.2	Name, title and experience of the quality management representative	16

3.8.3	Name and title of signatories for certification	16

3.8.4	Other employees in ExCB activity	17

3.9	Organizational structure	17

3.10	Indemnity insurance	17

3.11	Resources	17

3.12	Committees (such as governing or advisory boards)	17

3.13	Certification operations	18

3.13.1	National approval/certification methods	18

3.13.2	Certification policy	18

3.13.3	Application for certification	18

3.13.4	Certification decision	18

3.13.5	Suspension and cancellation of certificates	18

3.14	Certificates issued	18

3.15	National accreditation	20

3.16	Assessment of manufacturers and issue of QARs	20

3.17	Comments (including issues found during assessment)	20

ExTL for IECEx Certified Equipment Scheme	22

4.1	Assessment references	22

4.1.1	General references	22

4.1.2	Additional references applied for this assessment	22

4.2	Candidate ExTL persons interviewed	22

4.3	Associated ExCB(s)	22

4.4	Organisation	22

4.4.1	Names, titles and experience of the senior executives	23

4.4.2	Name, title and experience of the quality management representative	23

4.4.3	Other employees in ExTL activity	23

4.5	Organizational structure	24

4.6	Resources	24

4.7	Test reports issued	24

4.8	National accreditation	24

4.9	Calibration	25

4.10	Tests witnessed during the assessment visit	25

4.11	Participation in IECEx Proficiency Testing Programs	25

4.12	Comments (including issues found during assessment)	26

ExCB for Certified Service Facilities Scheme	27

5.1	Assessment references	27

5.2	Candidate ExCB persons interviewed	27

5.3	Organisation	27

5.3.1	Names, titles and experience of the senior executives	27

5.3.2	Name, title and experience of the quality management representative	27

5.3.3	Name and title of signatories for certification	28

5.3.4	Other employees in ExCB activity	28

5.4	Organizational Structure	28

5.5	Resources	28

5.6	Certification operations	28

5.6.1	National approval/certification Methods	28

5.6.2	Application for certification	28

5.6.3	Withdrawal and cancellation of certificates	28

5.7	Statistics	28

5.8	National accreditation	29

5.9	Comments (including issues found during assessment)	29

Annexes	30

Annex A Overall Organisation Chart	31

Annex B Organisation Chart of ExCB and ExTL	32

Annex C Accreditation Certificate for ISO/IEC 17065	34

Annex D Accreditation Certificate for ISO/IEC 17025	36

# Assessment information

## Type of Body covered by this assessment:

| ExCB for IECEx Certified Equipment Scheme |  |
| --- | --- |
| ExTL for IECEx Certified Equipment Scheme |  |
| ExCB for IECEx Certified Service Facilities Scheme |  |
| ExCB for IECEx Conformity Mark Licensing System |  |

NOTE 1	ExCB - IECEx Certification Body

NOTE 2	ExTL - IECEx Testing Laboratory

## Type of assessment:

| Pre-assessment for candidate body |  |
| --- | --- |
| Initial assessment for candidate body |  |
| Surveillance |  |
| Re-assessment |  |
| Scope extension |  |

## Details of body

### Country

Germany

### Name of body

DEKRA Testing and Certification GmbH

### Name and title of nominated principal contact

| Name | Title | E-mail address |
| --- | --- | --- |
| Dr Franz Eickhoff | Head of Division Explosion Protection Electrotechnology | franz.eickhoff@dekra.com |

## Assessment information

### Members of the assessment team

| Name | Role |
| --- | --- |
| Katy A. Holdredge | IECEx Lead Assessor |
| Julien Gauthier | IECEx Assessor |

### Place(s) of assessment

| Dinnendahlstrasse 9, 44809 Bochum, Germany | Adlerstrasse 29, 45307 Essen, Germany (Gas detection only) |
| --- | --- |

### Assessment date(s)

November 5 – November 7, 2019

## Application information and background information on the assessment

Information relevant to the Secretariat review process: ExMC/251B/Q ExCB/ExTL Capability Declaration for Extension of Scope from the Secretariat.

## Scopes

### ExCB scope for equipment certification scheme

| Number | Title | Comments, e.g., if scope change |
| --- | --- | --- |
| IEC 60079-0 Edition 7.0 | Explosive atmospheres - Part 0: Equipment - General requirements |  |
| IEC 60079-1 Edition 7.0 | Explosive atmospheres - Part 1: Equipment protection by flameproof enclosures “d” |  |
| IEC 60079-2 Edition 6.0 | Explosive atmospheres - Part 2: Equipment protection by pressurized enclosure «p» |  |
| IEC 60079-5 Edition 4.0 | Explosive atmospheres - Part 5: Equipment protection by powder filling «q» |  |
| IEC 60079-6 Edition 4.0 | Explosive atmospheres - Part 6: Equipment protection by oil immersion «o» |  |
| IEC 60079-7 Edition 5.1 | Explosive atmospheres - Part 7: Equipment protection by increased safety "e" |  |
| IEC 60079-11 Edition 6.0 | Explosive atmospheres - Part 11: Equipment protection by intrinsic safety “i” |  |
| IEC 60079-13 Edition 1.0 | Explosive atmospheres - Part 13: Equipment protection by pressurized room "p" and artificially ventilated room "v" |  |
| IEC 60079-15 Edition 5.0 | Explosive atmospheres – Part 15: Equipment protection by type of protection "n" |  |
| IEC 60079-16 Edition 1 | Explosive atmospheres – Part 16: Artificial ventilation for the protection of analyser (s) houses | See §4.8. |
| IEC 60079-18 Edition 4.1 | Explosive atmospheres – Part 18: Equipment protection by encapsulation “m” |  |
| IEC 60079-25 Edition 2.0 | Explosive atmospheres – Part 25: Intrinsically safe electrical systems |  |
| IEC 60079-26 Edition 3.0 | Explosive atmospheres - Part 26: Equipment with equipment protection level (EPL) Ga |  |
| *IEC 60079-27 Edition 2.0 | Explosive atmospheres – Part 27: Fieldbus intrinsically safe concept (FISCO) |  |
| IEC 60079-28 Edition 2.0 | Explosive atmospheres - Part 28: Protection of equipment and transmission systems using optical radiation |  |
| IEC 60079-29-1 Edition 2.0 | Explosive atmospheres - Part 29-1: Gas detectors – Performance requirements of detectors for flammable gases |  |
| IEC/IEEE 60079-30-1 Edition 1.0 | Explosive atmospheres – Part 30-1: Electrical resistance trace heating – General and testing requirements |  |
| IEC 60079-31 Edition 2.0 | Explosive atmospheres – Part 31: Equipment dust ignition protection by enclosure "t" |  |
| IEC 60079-32-2 Edition 1.0 | Explosive atmospheres - Part 32-2: Electrostatics hazards - Tests (may be used for testing purposes but not for issuing an IECEx Certificate of Conformity) |  |
| IEC 60079-33 Edition 1.0 | Explosive atmospheres – Part 33: Equipment protection by special protection “s” |  |
| IEC 60079-35-1 Edition 1.0 | Explosive atmospheres – Part 35-1: Caplights for use in mines susceptible to firedamp – General requirements – Construction and testing in relation to the risk of explosion | See §4.8. |
| IEC TS 60079-39 Edition 1.0 | Explosive atmospheres-Part 39: Intrinsically safe systems with electronically controlled spark duration limitation |  |
| IEC TS 60079-40 Edition 1.0 | Explosive atmospheres-Part 40: Requirements for process sealing between flammable process fluids and electrical systems | See §4.8. |
| IEC TS 60079-42 Edition 1.0 | Explosive atmospheres- Electrical safety devices for the control of potential ignition sources for Ex-Equipment | Scope increase |
| *DS 2015/001A | IECEx Assessment and Certification of Equipment assemblies | ExTAG DS withdrawn on 2020-01-01 |
| IEC TS 60079-46 Edition 1.0 | Explosive atmospheres - Part 46: Equipment assemblies |  |
| *IEC 61241-0 Edition 1.0 | Electrical apparatus for use in the presence of combustible dust - Part 0: General requirements |  |
| *IEC 61241-1 Edition 1.0 | Electrical apparatus for use in the presence of combustible dust - Part 1: Protection by enclosure “tD” |  |
| *IEC 61241-1-1 Edition 2 | Part 1-1: Electrical apparatus protected by enclosures and surface temperature limitation - Specification for apparatus |  |
| *IEC 61241-4 Edition 1.0 | Electrical apparatus for use in the presence of combustible dust - Part 4: Protection by pressurization "pD" |  |
| *IEC 61241-11 Edition 1.0 | Electrical apparatus for use in the presence of combustible dust – Part 11: Protection by intrinsic safety 'iD' |  |
| *IEC 61241-18 Edition 1.0 | Electrical apparatus for use in the presence of combustible dust - Part 18: Protection by encapsulation "mD" |  |
| IEC 61779-1 Edition 1 | Part 1: General requirements and test methods - Electrical apparatus for the detection and measurement of flammable gases |  |
| IEC 61779-2 Edition 1 | Part 2: Performance requirements for group I apparatus indicating a volume fraction up to 5% methane in air |  |
| IEC 61779-3 Edition 1 | Part 3: Performance requirements for group I apparatus indicating a volume fraction up to 100% methane in air |  |
| IEC 61779-4 Edition 1 | Part 4: Performance requirements for group II apparatus indicating a volume fraction up to 100% lower explosive limit |  |
| IEC 61779-5 Edition 1 | Part 5: Performance requirements for group II apparatus indicating a volume fraction up to 100% gas |  |
| *IEC 62013-1 Edition 2.0 | Caplights for use in mines susceptible to firedamp - Part 1: General requirements - Construction and testing in relation to the risk of explosion |  |
| *IEC 62086-1 Edition 1 | General and testing requirements - Electrical apparatus for explosive gas atmospheres – Electrical resistance trace heating |  |
| ISO 80079-36 Edition 1 | Explosive atmospheres - Part 36: Non-electrical equipment for explosive atmospheres - Basic method and requirements |  |
| ISO 80079-37 Edition 1 | Explosive atmospheres - Part 37: Nonelectrical type of protection constructional safety "c", control of ignition source "b", liquid immersion "k" |  |

NOTE 1	Standards shown with an asterisk (*) are superseded standards

### ExTL scope

The ExTL scope is the same as for the ExCB.

### ExCB scope for Service Facilities Scheme

DEKRA increased their scope for the service facilities scheme as part of this reassessment to include dusts, “t”, “tD” and “DIP”.

| 60079-19 Edition 3 | Explosive atmospheres – Part 19: Equipment repair, overhaul and reclamation |
| --- | --- |
| Plus, the following types of protection | ID |
| Flameproof Enclosure "d" | d |
| Increased Safety "e" | e |
| Type of Protection "n" | n |
| Intrinsic Safety "i" | i |
| Oil Filled "o" | o |
| Pressurisation "p" | p |
| Dusts to IEC 60079-31 "t" | t |
| Dusts to IEC 61241-1 "tD" | tD |
| Dusts to IEC 61241-1-1 "DIP" | DIP |
| Other | - |

# 2. Common information

## Legal entity of body

DEKRA Testing and Certification GmbH (DTC) has a legal entity from the “Handelsregister B des Amtsgerichts Stuttgart” under the number HRB 759624.  The legal entity address is in Stuttgart with the certification body and testing bodies addresses in Bochum and Essen, as shown on ISO/IEC 17065 and ISO/IEC 17025 accreditations.  See Annexes C and D.

The latest document, dated July 25th, 2019, was reviewed during the assessment and found to meet the requirements of the IECEx.

## Financial support

DEKRA Testing and Certification GmbH has its income from testing, certification, inspection and training.

## History

1894	Founded as BVS, Mine safety testing body

Since 1894	Investigations into the origin of explosions gas and dust explosions especially in hard coal mining underground

1912	Starting with certification of explosion proof equipment for coal mines and other industries

~ 1920	Transfer of research results into German rules and standards

1950ies	Development from research institute to approved testing body and service provider

1980	Accreditation under European Directives for explosion proof equipment for coal mines and other industries

1999	IECEx-TL & CB

2006 	Acknowledgement as Approved Inspection Body for explosion protection according to ATEX-Directive 99/92/EC

Since 2007 	Member of DEKRA group

Since 2019	Fusion with DEKRA Testing and Certification GmbH, Stuttgart (no changes in ExCB and ExTL).

DTC is a Notified Body in European Certification System. The Notified Body (certification body) number is 0158 for ATEX Directive 2014/34/EU. DTC is an accepted inspection body according to ISO/IEC 17020. See NANDO for further information on notification.

## Documentation

### Quality manual

The DEKRA Testing and Certification GmbH quality system is described in the Quality Manual, which covers the overall organization of DEKRA Testing and Certification GmbH including the Certification Body, Test Laboratory and the Inspection Body, issue date 2019-09-25). It is valid for all locations (Bochum: whole IECEx scope without gas detection; Essen: gas detection). It consists of a general part, the appendixes/procedures (A) and the forms (FRM).

DEKRA’s quality documentation was in transition during the assessment from the old system, available on their Intranet, to the new system, Sharepoint.

The Quality Manual was reviewed during the assessment and found to meet the requirements of the IECEx.

### Procedures

Procedures are available and controlled for the three main divisions dealing with IECEx using the following designations, ExE – electrical, GW – gas detection, and ExA – non-electrical.

The procedures reviewed during the assessment were found to meet the requirements of the IECEx.

### Work instructions

DEKRA Testing and Certification GmbH has two high-level work instructions one for the testing laboratory “Q-PB-25 Prüflabor”, dated 2019-09-30, and one for the certification body, “Q-PB-65 Zertifizierung”, dated 2019-09-09.

There were missing work instructions found during the audit relating to Witness testing per IECEx OD 024 and commenting on ExTAG DS.  They were subsequently revised and found to meet the requirements of the IECEx.

### Records (including test records where relevant)

All records are maintained and stored and are currently a mixture of hard copy and electronic records. The archiving process in place for all records is described in “QMH-05 Chapter 5.4 Control of Records”, dated 2019-09-12.  The document retention requirements in IECEx OD 207 are met.

### Document change control

Document change control is described in “QMH-05 Chapter 5.3 Control of Documents”, dated 2019-09-12.  It requires that three people be involved in all changes to the quality manual and two people for all other documentation, with a requirement for a yearly check for accuracy/validity.

This document was reviewed during the assessment and found to meet the requirements of the IECEx.

## Confidentiality

(For staff, contractors and members of advisory bodies)

The procedure is described in the Quality Management Handbook (QMH-01 chapter 1.2 Confidentiality, dated 2019-09-02).

Each employee signs a document concerning confidentiality as part of their initial contract. In addition, all staff sign the document Q-F-0, Commitment to impartiality, independent and confidentiality.

There is a separate contract for subcontracted quality auditors, FRM-90.10.01, Statement Auditor.  Two examples were viewed.

The procedure and some of the documents were checked and found to meet the requirements of the IECEx.

## Communication with public and customers (Hard copy and Electronic)

There is basic information available on DEKRA’s website that they can offer IECEx services: https://dekra-testing-and-certification.de/en/.

## Recognitions and agreements

DEKRA holds mutual recognition agreements with PTB (Germany), NEPSI (P.R.C.), ERTL East (India), KGS (South Korea), NANIO (Russia) and Mining Administration (China). Acceptance is granted by MSHA (USA) and United States Coast Guard (USCG).

## Internal audit

Internal audits are performed according to the Quality Management Handbook, chapter 5.3 (Q-PB-0 Interne Audits, dated 2019-09-03).

The internal audit from 2018 included the witness of two tests by a senior technical staff member, including Clause 10.12 of IEC 60079-11, 10.12, with no issues noted.  The remainder of the audit was conducted by the quality manager on 06.12.2018 and covered both ISO/IEC 17065 and ISO/IEC 17025, no non-conformities were noted.   An additional internal audit was conducted on 2 and 3 of May 2018, with no nonconformities related to IECEx.

The internal audit plan for 2019 was also checked and everything was found to meet the requirements of the IECEx.

## Management review

The management review is performed according to the Quality Management Handbook, chapter 5.8 and is performed one time per year. The last management review was held on 2019-04-04 and was completed by top management using the input of appropriate staff.

The minutes of the 2018 management review were presented during the assessment and found to meet the requirements of the IECEx.

## Contracting, subcontracting and witness testing

### Contracting

The issue of contracting/subcontracting is described in the QMH-02 chapter 2.7 Subcontracting, dated 2019-09-02.  DEKRA currently has five external assessors qualified to perform IECEx QARs with one more in process.

### Subcontracting

Subcontractors performing work related to IECEx are listed in A-300_11_01_03 of the Electrical division test laboratory. This document includes information on accreditation of the bodies.  None of the tests including in subcontracting are listed in the TCD as requiring minimum test equipment.  The divisions Gas Detection and Non-Electrical do not have any subcontractors for IECEx work.

The following tests are, or may be, subcontracted by the body:

| Standard | Clause | Test |
| --- | --- | --- |
| IEC 60079-7 | 6.6.3 | Shock test |
| IEC 60079-0 | 26.5 | Thermal tests on large electrical engines |
| IEC 60079-7 | 6.2 | Rotating electrical machines |
| IEC 60079-15 | 22.13 | Additional ignition tests for large or high-voltage machines |
| IEC 60079-28 | 6 | Optical ignition tests |
| IEC 60079-0 (ISO 4892-2) | 26.10 | UV resistance test |
| IEC 60079-30-1 | 5.1.9, 5.1.10, 5.1.11, 5.1.12, 5.1.13, 5.1.14, 5.1.15 and 5.1.16 | Tests on heat tracing cable |

### Witness testing

DEKRA has two levels of assessments, one for manufacturers that are assessed each time with all testing fully witnessed and one for manufacturers that perform tests more frequently and undergo annual assessments.  The form used for IECEx OD 024 assessments is in their document control system.  Recent records are stored on their Sharepoint drive and older records are stored hard copy. Two reports for the annual assessments of an active client and the contract were reviewed and were in compliance with the requirements of the IECEx.

There was no evidence of a written procedure for witness testing in accordance with IECEx OD 024 during the assessment and the registration of the facilities has not been done with the secretariat.  An existing procedure was subsequently revised and found to meet the requirements of the IECEx.

## Training and competence

This is covered in Clause 3.1, QMH-03.  All staff employed are selected for qualifications and/or experience relevant to their responsibilities. Each member of staff has a full job description, which comprehensively defines their responsibilities, job function, qualification requirements and their position within the organisation.

On regular basis there is training of people in the ExCB and ExTL on the operations, outcome of audits, revised standards and procedures related to IECEx. An example of a comprehensive training session presentation was shown and was found to meet the requirements of the IECEx. In addition, there is a training plan for the division in A-300.04.03.

There are competency matrices for ExCB and ExTL, which are included in the site assessment report.

The documentation was found to meet the requirements of the IECEx.

## Complaints and appeals (including appeals to IECEx)

There is a general process in DEKRA Testing and Certification GmbH for internal complaints, internal and external audits, and external complaints (QMH 4.7, dated 2019-09-02, “Q-PB-0 Umgang mit Beschwerden und Einsprüchen”, dated 2019-08-29).

There was no reference to the appeals process as described in IECEx 02 at the time of the assessment. The document was subsequently revised and found to meet the requirements of the IECEx.

## Impartiality

See as reference QMH 1.1 Impartiality, dated 2019-09-02.  There is a meeting once a year to discuss this issue, among others, involving a committee.  The last meeting was on May 9, 2019 and involved a variety of stakeholders.  DEKRA has an Excel spreadsheet, Q-A-0, that tracks all risks to impartiality.  The documentation was found to meet the requirements of the IECEx.

## Commenting on ExTAG Documents

There was no documented procedure/work instruction at the time of the assessment, but evidence of providing input was provided.  An update work instruction was subsequently provided and found to meet the requirements of the IECEx.

## Special facts to be noted

None.

## Supporting documentation

Copies of additional supporting information for this assessment have been provided to the applicant and the IECEx Secretariat.  These are included in a site assessment report or provided separately and include:

Details of issues raised and how these have been resolved, provided in F-004;

Checklist for ISO/IEC 17065;

Checklist for ISO/IEC 17025;

Completed Technical Capability Document (TCD); and

Photos of the facilities/tests witnessed are included in the above TCD.

## Recommendations

Based on the assessment performed on November 5 – November 7, 2019, DEKRA Testing and Certification GmbH is recommended for continued acceptance in the IECEx scheme as:

An ExCB in the IECEx Certified Equipment Scheme;

An ExTL in the IECEx Certified Equipment Scheme; and

An ExCB in the IECEx Certified Service Facilities Scheme.

This is according to the scope of the standards listed in this document, including the extension of scope, subject to resolution of the issues found during the assessment.

| Katy A. Holdredge | Julien Gauthier |
| --- | --- |
| IECEx Lead Assessor | IECEx Assessor |

Date:  October 14, 2020

# ExCB for IECEx Certified Equipment Scheme

## Assessment references

### General references

IECEx 02 IECEx Certified Equipment Scheme covering equipment for use in explosive atmospheres – Rules of Procedure

OD003-2  Assessment, surveillance assessment and re-assessment of ExCBs and ExTLs operating in the IECEx 02, IECEx Certified Equipment Scheme

ISO/IEC 80079-34 Edition 2, Explosive atmospheres – Part 34: Application of quality systems for equipment manufacture

OD009 Issuing of CoCs, ExTRs and QARs

IECEx Document OD 025 Guidelines on the Management of Assessment and Surveillance programs for the assessment of Manufacturer’s Quality Systems in accordance with the IECEx Scheme

OD026 IECEx Certified Equipment Scheme – Guidelines for the qualification of Lead Auditor and Auditors, in accordance with the IECEx System

ISO/IEC 17065: 2012, Edition 1, General requirements for bodies operating product certification systems Conformity assessment — Requirements for bodies certifying products, processes and services

IECEx Technical Capability Document (TCD)

ExTAG decision sheets (DSs)

NOTE	The latest editions of the above documents were applied

### Additional references applied for this assessment

- OD233 IECEx Certified Equipment Scheme - Assessment of Ex “s" Equipment

- OD280 IECEx Certified Equipment Scheme – Guide to Certification of Non-electrical Equipment and Protective Systems

## Candidate ExCB persons interviewed

| Name | Position |
| --- | --- |
| Dr Franz Eickhoff | Head of Division Explosion Protection Electrotechnology |
| Jörg Koch | Head of Certification Body |
| Dr. Michael Wittler | Deputy Head of Laboratory Deputy Head of Division Explosion Protection Electro Technology |
| Ralf Leiendecker | Deputy Head of specialist division |

## Associated ExTL(s)

The ExTL is integral with the ExCB.

The main laboratories for IECEx testing are located at Dinnendahlstrasse 9, 44809 Bochum, Germany.

## Associated certification functions

DEKRA Testing and Certification provides ATEX certificates in accordance with ATEX Directive 2014/34/EU and their operations are accredited to IEC/ISO 17065 by the National Accrediting body DAkkS.

## National marks and certificates

DEKRA Testing and Certification GmbH issued ATEX certificates are accepted by the European commission and by European state regulators.

Other divisions of the company issue additional certificates in scope of EU or German legislation.

## Standards accepted

See clause 1.6 of this report

## National differences to IEC standards

National differences to IEC standards are those for the German differences listed in the latest version of the IECEx Scheme Bulletin.

## Organisation

### Names, titles and experience of the senior executives

| Name | Title | Experience |
| --- | --- | --- |
| Jörg-Timm Kilisch | CEO | - |
| Dr. Jörg Kiesewetter | Deputy Head of Certification Body Deputy Head of Laboratory Deputy Head of Division Personal Protective Equipment | - 10 years |
| Dr. Michael Wittler | Deputy Head of Laboratory Deputy Head of Division Explosion Protection Electro Technology | - 10 years |
| Dr. Michael Sippel | Deputy Head of Division Explosion Protection Plant Safety Deputy Head of Inspection Body | - 10 years |
| Dr. Franz Eickhoff | Deputy Head of Certification Body Head of Division Explosion Protection Electro Technology | - 10 years |

### Name, title and experience of the quality management representative

| Name | Title | Experience |
| --- | --- | --- |
| Dr. techn. Nina Thimm | Head of Quality Management | - 10 years |
| Jörg Koch | Head of Certification Body | - 10 years |
| Daniel Horn | Quality Manager Division Explosion Protection Electro technology | - 6 years |

### Name and title of signatories for certification

| Name | Title |  | Comments |
| --- | --- | --- | --- |
| Jörg Koch | Head of Certification Body |  | - 10 years |
| Dr.-Ing. Michael Wittler |  | Head of specialist division Deputy Head of Laboratory Deputy Head of Division Explosion Protection Electro Technology | - 10 years |
| Dr.-Ing. Christiane Sultan |  | Deputy Head of specialist division | - 10 years |
| Ralf Leiendecker |  | Deputy Head of specialist division | - 10 years |
| Stefan Wiegand |  | Deputy Head of Certification Body Head of specialist division Head of Division Safety of Machinery and Components | - 10 years |
| Dr. Franz Eickhoff |  | Deputy Head of Certification Body Head of Division Explosion Protection Electro Technology | - 10 years |

### Other employees in ExCB activity

| Name | Title | Responsibility and Experience in Ex |
| --- | --- | --- |
| Stefanie Arndt | Office member | - |
| Silke Knitter | Office member | - |
| Dagmar Muskatewitz | Office member | - |
| Christiane Pohl | Office member | - |
| Brigitte Szczesny | Office member | - |

## Organizational structure

See organisation charts in Annex B.

## Indemnity insurance

DEKRA Testing and Certification GmbH holds professional indemnity and public liability insurance. These are covered in one policy from Gothaer Allgemeine Versicherungs AG in Cologne, Germany. The insurance number is 38.660.561932 DEL000396180.

The coverage is considered acceptable.

## Resources

DEKRA Testing and Certification appears to have an adequate number of staff for the current level of business. The laboratory and offices are located in an industrial unit which provides an adequate environment for the work.

The testing equipment is suitable for the range of tests carried out in house.

## Committees (such as governing or advisory boards)

The composition and terms of reference of the Advisory Board are described in the Quality Management Handbook “Chapter 1.1.2 Mechanismus zur Sicherung der Unparteilichkeit”, Q-A-65 Mechanismus zur Sicherung der Unparteilichkeit”, dated 2019-10-31. The Ex Committee comprises representatives of market surveillance, research and education and industry representatives with no single interest predominating.

The content of the procedure meets the requirements of ISO/IEC 17065 and the IECEx requirements.

## Certification operations

### National approval/certification methods

DEKRA Testing and Certification is recognised under the National accreditation systems and schemes. It has procedures for compliance with IECEx Rules and Operational Documents.

### Certification policy

Statement available in Quality Management handbook (see reference Q-A-0 Qualitätspolitik, dated 2019-09-09) and further applicable documents.

NOTE Typically this may be a separate policy or included in the quality policy

### Application for certification

Procedure available in Quality Management handbook (see reference Q-PB-65 Zertifizierung, Chapter 3 Application, dated 2019-09-09) and further applicable documents.

### Certification decision

Procedure available in Quality Management handbook (see reference Q-PB-65 Zertifizierung, Chapter 7 Certification decision, dated 2019-09-09) and further applicable documents.

### Suspension and cancellation of certificates

Procedure available in Quality Management handbook (see reference Q-PB-65 Zertifizierung, Chapter 12 Termination, reduction, suspension or withdraw, dated 2019-09-09) and further applicable documents (etc. “Besondere Bedingungen FRM-90.103.08”).

At the time of the assessment DEKRA had one IECEx certificate that is suspended and others that have been cancelled.

## Certificates issued

Number of certificates issued under for the preceding four years for each type of protection.

| Standard numbers | Type of protection or other identifying information | Number of issued certificates (for last 4 years) |  |  |  | Total (only IECEx) |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | 2015 | 2016 | 2017 | 2018 |  |
| IEC 60079-0 | General requirements |  |  |  |  | Parts 0 included in numbers below |
| IEC 60079-1 | Flameproof enclosures 'd' | 186 (75) | 188 (73) | 186 (88) | 155 (83) | 715 (319) |
| IEC 60079-2 | Pressurized enclosure 'p' | 51 (12) | 35 (4) | 26 (9) | 37 (9) | 149 (34) |
| IEC 60079-5 | Powder filling 'q' | 11 (3) | 7 (3) | 11 (6) | 19 (5) | 48 (17) |
| IEC 60079-6 | Oil immersion 'o' | 5 (0) | 3 (1) | 1 (0) | 6 (0) | 15 (1) |
| IEC 60079-7 | Increased safety 'e' | 201 (90) | 144 (81) | 170 (88) | 202 (98) | 717 (357) |
| IEC 60079-11 | Intrinsic safety 'i' | 304 (113) | 289 (118) | 273 (109) | 212 (110) | 1078 (450) |
| IEC 60079-13 | Equipment protection by pressurized room 'p' | - | 1 (1) | - | - | 1 (1) |
| IEC 60079-15 | Type of protection 'n' electrical apparatus | 117 (48) | 86 (46) | 75 (35) | 53 (22) | 331 (151) |
| IEC 60079-16 | Artificial Ventilation for the Protection of Analyzer(s) Houses | - | - | - | - | 0 (0) |
| IEC 60079-18 | Encapsulation 'm' electrical equipment | 37 (17) | 26 (8) | 25 (13) | 47 (16) | 135 (54) |
| IEC 60079-25 | Intrinsically safe systems | 1 (0) | 2 (1) | 6 (3) | 1 (0) | 10 (4) |
| IEC 60079-26 | Equipment with equipment protection level (EPL) Ga | 32 (16) | 34 (17) | 24 (12) | 40 (20) | 130 (65) |
| IEC 60079-27 | Fieldbus intrinsically safe concept (FISCO) and Fieldbus non-incendive concept (FNICO) | - | - | - | - | 0 (0) |
| IEC 60079-28 | Optical radiation | 6 (3) | 31 (4) | 25 (14) | 27 (11) | 89 (32) |
| IEC 60079-29-1 | Gas detectors | 12 (0) | 9 (0) | 7 (0) | 12 (0) | 40 (0) |
| IEC 60079-30-1 | Electrical resistance trace heating | 0 (0) | 0 (0) | 0 (0) | 1 (1) | 1 (1) |
| IEC 60079-31 | Protection by enclosure ‘t’ | 223 (96) | 220 (99) | 202 (95) | 178 (92) | 823 (382) |
| IEC/IEEE 60079-30-1 | Electrical resistance trace heating - General and testing requirements | 0 (0) | 0 (0) | 0 (0) | 1 (1) | 1 (1) |
| 60079-33 | Explosive atmospheres – Part 33: Equipment protection by special protection “s” | 0 (0) | 0 (0) | 0 (2) | 0 (1) | 0 (3) |
| 60079-35-1 | Caplights for use in mines susceptible to firedamp – General requirements – Construction and testing in relation to the risk of explosion | -- | -- | -- | -- | -- |
| TS 60079-39 | Intrinsically safe systems with electronically controlled spark duration limitation | -- | -- | -- | -- | -- |
| TS 60079-40 | Requirements for process sealing between flammable process fluids and electrical systems | -- | -- | -- | -- | -- |
| TS 60079-42 | Electrical safety devices for the control of potential ignition sources for Ex-Equipment | -- | -- | -- | -- | -- |
| EN 50495 | Safety devices required for the safe functioning of equipment with respect to explosion risks | 3 | 4 | 4 | 3 | 14, together with EN ISO 13849 or IEC 61508 (0) |
| EN ISO 13849-1 / IEC 61508 | Safety of machinery - Safety related parts of control systems, Part 1: General principles for design / Functional safety of electrical / electronic / programmable electronic safety-related systems - Part 1: General requirements | 3 (0) | 4 (0) | 4 (3) | 3 (4) | 14 (7) |
| TS 60079-46 | Equipment assemblies | 0 (0) | 0 (0) | 0 (0) | 0 (0) | 0 (0) |
| EN 13463-1 | Non-electrical equipment for use in potentially explosive atmospheres - Part 1: Basic method and requirements | 102 (0) | 108 (0) | 71 (0) | 71 (0) | 352(0) |
| EN 13463-5 | Non-electrical equipment intended for use in potentially explosive atmospheres - Part 5: Protection by constructional safety 'c' |  |  |  |  |  |
| EN 13463-6 | Non-electrical equipment for use in potentially explosive atmospheres - Part 6: Protection by control of ignition source "b" |  |  |  |  |  |
| EN 13463-8 | Non-electrical equipment for potentially explosive atmospheres - Part 8: Protection by liquid immersion 'k' |  |  |  |  |  |
| ISO 80079-36 | Non-electrical equipment for explosive atmospheres – Basic method and requirements |  |  |  |  | Certification started in 2016-08-12 |
| ISO 80079-37 | Non-electrical equipment for explosive atmospheres – Non electrical type of protection constructional safety “c”, control of ignition source “b”, liquid immersion “k” |  |  |  |  |  |

NOTE	Above include certificates to IEC 60079-0 unless otherwise shown

## National accreditation

See reference in Annex C.

NOTE The national accreditation is checked annually by the IECEx Secretariat.

## Assessment of manufacturers and issue of QARs

This is documented FRM-90.179.01 covers the requirements for auditor competence.  There is a form that is completed when initially qualified FRM-90.05.04 and for monitoring FRM-90.08.01.  The documentation was viewed for a particular employee.  The assessment of manufacturers is done in different scopes:

- ISO 9001:2015: as subcontractor for DEKRA Certification GmbH and TÜV Nord Systems GmbH;

- ATEX Directive with ISO/IEC 80079-34 Ed. 2 as Notified body No. 0158; and

- IECEx Scheme with ISO/IEC 80079-34 Ed. 2 as Certification Body.

- Audit report (Q-F-21), NCR (FRM-90.16.04 or Q-F-21-ExE-Abweichungsdokumentation), form for people who were in the audit available on Microsoft Access Project database.

- A-300_10 Arbeitsblatt – contains list of documents needed for the audit.

During the time of the assessment there was no evidence that the management of QARs and CoCs linked to other ExCB’s QARs is managed as required by IECEx OD 250, especially Clauses 1.2.2 and 1.2.3.  Procedures were revised and reviewed by the assessment team and found to meet IECEx requirements.

.

## Comments (including issues found during assessment)

The following issues found during the site assessment:

- No evidence of procedure or registration of the manufacturers for IECEx OD 024;

- No evidence that manufacturers have been informed about their right to appeal as described in IECEx 02;

- No written procedure for commenting on ExTAG decision sheets;

- Certification found that had been suspended for over 12 months

- No evidence of the management of QARs and CoCs linked to other ExCBs’ QARs;

- No evidence provided that a manufacturer was following the requirements of IECEx OD 203 for their local assemblers; and

- IECEx certification issued for equipment that was described as a safety device.

Detailed information on this is shown in the site assessment report.  All issues have been resolved to the satisfaction of the assessment team.

# ExTL for IECEx Certified Equipment Scheme

## Assessment references

### General references

- IECEx 02 IECEx Certified Equipment Scheme covering equipment for use in explosive atmospheres – Rules of Procedure

IECEx OD003-2 Assessment, surveillance assessment and re-assessment of ExCBs and ExTLs operating in the IECEx 02, IECEx Certified Equipment Scheme

IECEx OD009 Issuing of CoCs, ExTRs and QARs

ISO/IEC 17025:2017 Edition 3, General requirements for the competence of testing and calibration laboratories

IECEx Technical Capability Document (TCD)

ExTAG decision sheets (DSs)

OD 202 IECEx Certified Equipment Scheme – IECEx Proficiency Testing Program

NOTE	The latest editions of the above documents were applied.

### Additional references applied for this assessment

- OD233 IECEx Certified Equipment Scheme - Assessment of Ex “s" Equipment

- OD280 IECEx Certified Equipment Scheme – Guide to Certification of Non-electrical Equipment and Protective Systems

## Candidate ExTL persons interviewed

| Name | Position |
| --- | --- |
| Marco Altkemper | Technician |
| Manfred German | Technician |
| Alexander Repp (ESD) | Technician |
| Rüdiger Losch | Technician |
| Steffen Scholle | Expert |
| Ute Hauke | Expert |
| Philipp Rethemeier | Expert |
| Stephan Ruhnau | Expert |
| Michael Wittler | Deputy Head of Laboratory |
| Rudolf Hauke | Laboratory Manager |

## Associated ExCB(s)

Integral with ExCB of DEKRA Testing and Certification GmbH, Bochum, Germany.

## Organisation

See Annexes B and C.

### Names, titles and experience of the senior executives

| Name | Title | Experience |
| --- | --- | --- |
| Jörg-Timm Kilisch | CEO | - |
| Dr. Jörg Kiesewetter | Deputy Head of Certification Body Deputy Head of Laboratory Deputy Head of Division Personal Protective Equipment | - 10 years |
| Dr. Michael Wittler | Deputy Head of Laboratory Deputy Head of Division Explosion Protection Electro Technology | - 10 years |
| Dr. Michael Sippel | Deputy Head of Division Explosion Protection Plant Safety Deputy Head of Inspection Body | - 10 years |

### Name, title and experience of the quality management representative

| Name | Title | Experience |
| --- | --- | --- |
| Dr. techn. Nina Thimm | Head of Quality Management | 10 years |
| Jörg Koch | Head of Certification Body | 10 years |
| Daniel Horn | Quality Manager Division Explosion Protection Electro technology | 6 years |

### Other employees in ExTL activity

| Name | Title/responsibility | Experience in Ex |
| --- | --- | --- |
| Peter Althoff | Expert | - 10 years |
| Marco Altkemper | Technician | - 10 years |
| Dr. Gregor Arnold | Expert | - 10 years |
| Klaus Bendieck | Expert | - 5 years |
| Stamena Boka Batakoa | Expert | - 5 years |
| Mavin Droste | Technician | - 2 years |
| Dr Franz Eickhoff | Expert | - 10 years |
| Silviya Fromme | Expert | - 4 years |
| Manfred German | Technician | - 6 years |
| Rudolf Hauke | Expert | - 10 years |
| Ute Hauke | Expert | - 10 years |
| Olga Hildebrandt | Expert | - 5 years |
| Daniel Horn | Expert | - 4 years |
| Matthias Hurnicki | Expert | - 2 years |
| Thomas Kircher | Expert | - 10 years |
| Ralf Leiendecker | Expert | - 10 years |
| Rüdiger Losch | Technician | - 10 years |
| Oliver Marx | Technician | - 5 years |
| Bernd Melüh | Expert | - 7 years |
| Sebastian Pehl | Expert | - 6 years |
| Deniz Pezzutto | Expert | - 10 years |
| Alexander Repp (ESD) | Technician | - 5 years |
| Philipp Rethemeier | Expert | - 3 years |
| Frank Rippert | Expert | - 10 years |
| Stephan Ruhnau | Expert | - 10 years |
| Steffen Scholle | Expert | - 4 years |
| Dr. Christiane Sultan | Expert | - 5 years |
| Dr Michael Wittler | Expert | - 10 years |
| Felix Wloka | Expert | - 3 years |

## Organizational structure

See Annexes B and C.

## Resources

DEKRA has sufficient qualified staff, facilities and equipment to perform the testing within their scope.

## Test reports issued

The number of test reports issued is essentially equivalent to the number of certificates issued.  See 3.14.

## National accreditation

See Annex D.  The accreditation is currently to the 2005 edition of ISO/IEC 17025.  However, an audit was conducted by DAkkS on October 7 – 10, 2019 to update to the 2017 edition.

There are four standards within their scope as an ExTL that are not part of their ISO/IEC 17025 accreditation, IEC 60079-16, IEC 60079-35-1, IEC TS 60079-40 and IEC TS 60079-42.  It was verified that similar standards were in their scope, e.g., IEC 60079-13 for IEC 60079-16, IEC 61508 and EN 50495 for IEC TS 60079-42, or that they had the technical expertise to perform work to these standards as part of this assessment.  Their plan is to update their scope to include these standards within one year or drop them from their scope as an ExTL and ExCB.  The IECEx secretariat will follow-up on this activity as part of their normal annual check of accreditation.

DEKRA has a flexible scope through DAkkS, which allows them to include new editions of standards to their scope using their internal procedure, Q-PB-25, Revision 00, dated 29.08.2019 between their reassessments.  The quality manager retains an Excel spreadsheet of the scope, Listed Prufverfahren, including those added under the flexible scope option.  The procedure requires completion of FRM-00.35.01, dated 31.05.2017.  An example for Ed. 7.0 of IEC 60079-0 was sighted. There is QA documentation is on new Sharepoint website.

NOTE The national accreditation is checked annually by the IECEx Secretariat.

## Calibration

As reference see Q-PB-25 Prüfmittel, dated 2019-09-03.

Calibration and verification certificates were reviewed and found satisfactory for each test witnessed during the assessment. A minor issue was raised for one document and corrected during the assessment.

## Tests witnessed during the assessment visit

The following tests were witnessed during the assessment visit:

| Standard and edition | Clause number | Test | Comments |
| --- | --- | --- | --- |
| IEC 60079-0 Ed. 7.0 | 26.4.5 | Degree of protection (IP) by enclosures – dust test | Test performed correctly. |
| IEC 60079-0 Ed. 7.0 | 26.5.1 | Temperature measurement tests | Test performed correctly. Some minor issues were solved during the assessment. |
| IEC 60079-0 Ed. 7.0 | 26.13 | Surface resistance test of parts of parts of enclosures of non-metallic materials | Test performed correctly. |
| IEC 60079-1 Ed. 7.0 | 15.2.2 | Determination of Reference Pressure | Test performed correctly. |
| IEC 60079-11 Ed. 6.0 | 10.1 | Spark ignition test | Test performed correctly. Some minor issues were solved during the assessment. |
| IEC 60079-11 Ed. 6.0 | 10.5 | Tests for cells and batteries | Test performed correctly. |
| IEC 60079-18 Ed. 4.0 | 8.1.2 | Dielectric strength tests on compound | Test not fully performed, only satisfactory verification of availability of test equipment and staff experience by simulation of test. |
| IEC 60079-28 Ed. 2.0 | 5.2.2.2 | Optical power | Test performed correctly. |
| IEC 60079-29-1 Ed. 2.0 | 5.4.15 | Time of response | Test performed correctly. |
| IEC 60079-32-2 | 4.11 | Transferred charge | Test performed correctly. But the reference sample does not comply with standard requirement. |

All the tests were well prepared and performed in a timely manner. All people interviewed seem competent for the performance of the relevant test and test equipment were available.

## Participation in IECEx Proficiency Testing Programs

Program: PTB Ex PT Scheme, no other program

NOTE  It is anticipated that the assessment team will be provided with information on the performance of the body in completed PTB Ex PT Scheme programs.

| Year(s) of participation | IECEx Proficiency Testing program | General information about results |
| --- | --- | --- |
| 2011 | Pressure test | No warning or action signals |
| 2011 | Spark test | No warning or action signals |
| 2013 | Determination of temperature class | No warning or action signals |
| 2013 | Flame transmission | No warning or action signals |
| 2015 | Intrinsic safety | No warning or action signals |
| 2015 | Electrostatic charge | No warning or action signals |
| 2017 | Reference pressure | No warning or action signals |
| 2017 | Pressurised equipment | No warning or action signals |
| 2019 | Enclosure test | No results yet |
| 2019 | Battery test for intrinsic safety “I” | No results yet |

## Comments (including issues found during assessment)

- Equipment missing for IEC 60079-32-2 test;

- Internal work instructions allowed waiving of test required by the standard without supporting technical justification; and

- Test reports missing supporting temperature measurement, dielectric strength or other supporting technical data.

Detailed information on this is shown in the site assessment report.  All issues have been resolved to the satisfaction of the assessment team.

# ExCB for Certified Service Facilities Scheme

## Assessment references

- IECEx 03  IECEx Certified Service Facilities Scheme covering repair and overhaul of Ex equipment – Rules of Procedure

IECEx OD316-5 IECEx Operations Manual – IECEx Certified Service Facilities Scheme – Part 5: Repair, overhaul and reclamation of Ex equipment. Assessment procedures for IECEx acceptance of Candidate Certification Bodies (ExCBs) for the purpose of issuing IECEx Certificates to Ex Service Facilities involved in the repair, overhaul and reclamation of Ex equipment

ISO/IEC 17065: 2012, Edition 1, General requirements for bodies operating product certification systems Conformity assessment — Requirements for bodies certifying products, processes and services

IEC 60079-19 Explosive atmospheres – Part 19: Equipment repair, overhaul and reclamation

NOTE	The latest editions of the above documents were applied

## Candidate ExCB persons interviewed

| Name | Position |
| --- | --- |
| Thomas Kircher | Expert / Lead Auditor |

## Organisation

See Annex B.

### Names, titles and experience of the senior executives

| Name | Title | Experience |
| --- | --- | --- |
| Dr. Michael Wittler | Deputy Head of Laboratory Deputy Head of Division Explosion Protection Electro Technology | 10 years |
| Dr. Franz Eickhoff | Deputy Head of Certification Body Head of Division Explosion Protection Electro Technology | 10 years |

### Name, title and experience of the quality management representative

| Name | Title | Experience |
| --- | --- | --- |
| Dr. techn. Nina Thimm | Head of Quality Management | 10 years |
| Jörg Koch | Head of Certification Body | 10 years |
| Daniel Horn | Quality Manager Division Explosion Protection Electro technology | 6 years |

### Name and title of signatories for certification

| Name | Title | Comments |
| --- | --- | --- |
| Dr.-Ing. Michael Wittler | Head of specialist division Deputy Head of Laboratory Deputy Head of Division Explosion Protection Electro Technology | 10 years |
| Dr.-Ing. Christiane Sultan | Deputy Head of specialist division | 10 years |
| Ralf Leiendecker | Deputy Head of specialist division | 10 years |
| Dr. Franz Eickhoff | Deputy Head of Certification Body Head of Division Explosion Protection Electro Technology | 10 years |

### Other employees in ExCB activity

Detailed as part of the staff competency matrix referenced in F-004.

## Organizational Structure

See Annex B.

## Resources

DEKRA has sufficient resources for the Certified Service Facility Scheme.

## Certification operations

### National approval/certification Methods

DEKRA Testing and Certification is recognised under the National accreditation systems and schemes. It has procedures for compliance with IECEx Rules and Operational Documents.

### Application for certification

See 3.13.3.

### Withdrawal and cancellation of certificates

See 3.13.5.

## Statistics

Detail experience in assessment and certification of Ex related Service Facilities for the Ex Protection under this application during the past 2 years:

flameproof		d		1 company, one location every year

increased safety	e		1 company, one location every year

type of protection	n		1 company, one location every year

intrinsic safety		i		2 companies, four locations every year

oil immersion		o		--

pressurization		p		1 company, one location every year

apparatus for Dusts	t, tD and DIP	1 company, one location every year

other

m		1 company, one location every year

trace heating	1 company, one location every year

## National accreditation

ZÜS-Benennung, German special construction “Amtlich anerkannte befähigte Person”.

## Comments (including issues found during assessment)

There were no issues found related to the Certified Service Facility Scheme.

# Annexes

Overall Organisation Chart

Organisation Chart of ExCB and ExTL

Accreditation Certificate for ISO/IEC 17065

Accreditation Certificate for ISO/IEC 17025