COLLECTION OF IECEx / ExTAG DECISION

| Standard: IEC 60079-1 Edition 6 IEC 60079-1 Edition 5 | Clause: Clause 15.1.2 (for both editions) |  |
| --- | --- | --- |
| Subject: Determination of explosion pressure (reference pressure) Status of document: Approved | Key words: - Explosion pressure determination - Proficiency testing | - Date: 2012-9-18 Originator of proposal: T. Krause/PTB TC/SC involved: |
| Recommendation to improve the effectiveness of the determination of explosion pressure Background: PTB and ExTAGWG10 have recently conducted a pilot phase of proficiency testing for IECEx and this has included explosion pressure determination in accordance to IEC 60079-1. An immediate and positive outcome from the program has been that it has highlighted a number of aspects that ExTLs can apply to readily improve the consistency of results across ExTLs. Question: What are the steps that can be taken to improve the consistency of results obtained when doing explosion pressure determination testing in accordance to IEC 60079-1? Answer: There are a number of readily applied measures that can be applied to improve results with a view to improved consistency across ExTLs in the IECEx Equipment Certification Scheme, these include (but are not necessarily limited to): - Most importantly is a critical analysis and plausibility check of all measured results (including pressure curves). - To avoid temperature influences on pressure transducers, the transducer face (diaphragm) has to be temperature compensated. If pressure transducers are not innately temperature compensated the best approach is to use a thin film (about 1 mm) of silicon rubber (allowing this to set) and then a film of silicon grease (about 3 mm). Other appropriate, but not so durable, approaches include the use of only silicon grease or vinyl electrical tape. The color of the compensation materials has no relevant effect on the measured explosion pressure and is freely selectable. Background: Almost all pressure transducers are sensitive to thermal influences. The reason is that the heat of the explosion process can cause an expansion of the housing of the internal crystal of the diaphragm. That expansion causes a pressure discharge of the crystal and simulates a low explosion pressure. - Individual measuring ranges (for time base and amplitude) depending on the test sample and the gas-air mixture should be used. Predefined settings are inadvisable. The focus should be on recording the rise time and the maximum amplitude. Background: The precision of measurement depends on the sample rate, the memory depth and the measuring range of the data acquisition system. Measuring ranges which are too wide could cause less precision and low details of the measuring result. Because of this it could happen that the data acquisition system doesn‘t record the highest measuring value and important peaks could be missed. - When using digital oscilloscopes or other measuring devices, DC coupling should always be selected. Background: AC coupling places a capacitor in line to the measuring input of the measuring system. This works like a high pass filter and blocks DC and low frequency AC signals which can result in lower measured explosion pressures than expected. - Shut off valves of the test facility should be located close to the test sample. Background: The volume of the gas connection lines could influence the volume of the test sample which causes different pressure values than expected. - Pressure transducers should be tightened to torque recommended by the pressure transducer manufacturer. Background: Wrong torque could causes side loading of the transducers diaphragm which results in distortions in the output signal and different pressure values than expected. - Pressure transducer faces should be flush with the inside wall of the enclosures in which they are mounted. Background: Flush mounted transducer faces could avoid an acoustic oscillation and minimize turbulences. - Measures should be taken to avoid excessive heating (significantly exceeding the ambient temperature) of the gas-air mixture inside of a test sample during an explosion test series (as a cause of the heated test sample). This can be e.g. a sufficient time between ignitions or the monitoring of the gas-temperature. The best location to monitor the temperature of the gas-air mixture is inside the sample. If this is not practicable another appropriate location is at the outlet of the sample. The temperature sensor should be directly in contact with the gas air mixture. Background: The density of the gas-air mixture depends on temperature. The measured explosion pressure depends on density of the gas-air mixture and therefore it depends on temperature. The explosion pressure decreases when the temperature of the gas-air mixture rises e.g. as a cause of the heated test sample. - To prevent pre-pressure inside of the test sample a delay time for shut off valves after purging with the gas air mixture should be used. Background: If shut off valves close immediately after purging with the gas air mixture the pressure inside of the sample could be higher than atmospheric pressure. Pre-pressure of the gas air mixture inside of the sample could cause higher explosion pressures than expected. |  |  |